**IVIS QUICK USER GUIDE:**

1. Double-click the Living Image software icon on the desktop.

2. In the dialog box that appears, select a user ID from the drop-down list. If the user ID is password protected, enter the password and click OK.



3. **Initializing the Imaging System:** It will take several miniutes for the check up the system and for camera to coll down to -90 (can be check by clicking on the red bar at the bottom of control window) Once the system is initialized and the camera is at proper temperature, red bar will turn green.



Red color appears in the

Control panel during

Initialization. The color

Turns blue when

Initialization is finished.



White – System is not initialized.

Red – System is initialized, but camera temperature is out of range and not ready for imaging.

Green – System is initialized and camera is at or within acceptable rangeof the demand temperature and locked. The system is ready for imaging.

White – System is not initialized.

Red – System is initialized, but CCD camera

temperature is out of range and not ready for imaging.

Green – System is initialized and CCD camera is at or within acceptable range

of the demand temperature and locked. The system is ready for imaging.

4) **Start anesthesia/ isoflurane system**

 a. Turn ON the O2 supply at the tank. Make sure your tank has enough oxygen.

 b. Turn ON oxygen supply from the chamber.

 c. Follow instruction label and turn on all keys accordingly.

5) Place your mice for anesthesia in the induction chamber. Close it tightly. Let animals sleep.

6) **Acquisition of bioluminescence image:**

a. Put a check mark next to Luminescent and select Auto exposure.

b. Choose Photograph (selecting Photograph automatically selects Overlay).

d. Select a field of view “C”.

e. Hit acquire button.

f. Image label box will open up. Fill up image/animal information in it.

7) **Saving image:** The software will ask you if you want to autosave your images, this feature will save all images to the specified location. You can also manually save images. If an image is not saved, when you try to close the image or software a warning will indicate that you should save the image.